## Code
ATT/ELP/YY/XXXXX-TTT
ATT | mean | ALS Testing Services (Thailand) Co., Ltd.
ELP | mean | Electronic job billing by ALS Pathumthani
ELS | mean | Electronic job billing by ALS Singapore
FA | mean | Failure analysis
ELWA | mean | Samples are water and air
GRP | mean | Green test
YY | mean | Year 2 digit
XXXXX | mean | Running number of report 5 digit
TTT | mean type of test
D | mean | DHS Test
H | mean | HPA Test: HPA filtration, Tape test, swab
L | mean | LPC test
IB | mean | IC Test under BOI
F | mean | NVR, FTIR test
G | mean | GC/MS test: Organic residue, GC/MS extractable
COR | mean | Corrosion
COP | mean | Copper wire test

## Coverpage-TH
60
Converter Ver 1.3 Ready
CUSTOMER PO NO.: | [6 empty] | D3 | h | 2 | nb | ns | z
ALS THAILAND REF NO.: | R/ATT/ELP/13/0532-D | [5 empty] | D2 | c | 2 | nb | ns | z
OUR REF NO.: | [7 empty] | c | 2 | nb | ns | z
DATE: | 2013-09-06 | [6 empty] | c | 2 | nb | ns | z
COMPANY: | Enpro Products (Thailand) Co., Ltd. | [5 empty] | D5 | c | 2 | nb | ns | z
DATE SAMPLE RECEIVED: | 2013-02-14 | [5 empty] | D1 | c | 2 | nb | ns | z
DATE ANALYZED: | 2013-02-28 | [6 empty] | c | 2 | nb | ns | z
DATE TEST COMPLETED: | 2013-03-01 | [6 empty] | c | 2 | nb | ns | z
SAMPLE DESCRIPTION: | One lot of sample was received with references: | [6 empty] | h | 2 | nb | ns | z
| Description: Antitle Disk | [5 empty] | D9 | h | 2 | nb | ns | z
| Model: Tahoe | [5 empty] | D10 | c | 2 | nb | ns | z
| Surface Area: 14,074 sq.mm | [5 empty] | D11 | c | 2 | nb | ns | z
| Remark: 24 pcs | [5 empty] | D12 | c | 2 | nb | ns | z
METHOD/PROCEDURE: | [7 empty] | r | a
Test | Procedure No | Number of pieces used for extraction | Extraction Medium | Extraction Volume |  |  |  | h | 5 | b | s | z
IC | 2092-000121 Rev AB | 0 | Ultrapure water at 80°C | 0mL |  |  | M1M2 | c | 5 | b | ns | z
NVR/FTIR | 92-004132 Rev Prelim CD | 0 | n-Hexane (HPLC Grade) | mL |  |  | M3 | c | 5 | b | ns | z
NVR | 92-004132 Rev Prelim CD | 0 | n-Hexane (HPLC Grade) | mL |  |  | M4 | c | 5 | b | ns | z
FTIR | 92-004132 Rev Prelim CD | 0 | n-Hexane (HPLC Grade) | mL |  |  | M5 | c | 5 | b | ns | z
LPC | 92-004230 Rev AJ | 1 per run | 0.1um Filtered Degassed DI Water | 1mL |  |  | M6 | c | 5 | b | ns | z
CVR | 92-004262 Rev AB | 1 | 18hrs baking at 65°C 250ml I-Chem Jar 1.5ml Methylene Chloride | - |  |  | M8 | c | 5 | b | ns | z
Extractable Copper | 2092-000122 Rev AA | 0 | Ultrapure water at 80°C | mL |  |  | M9 | c | 5 | b | ns | z
Extractable Platinum | 2092-000122 Rev AA | 0 | Ultrapure water at 80°C | mL |  |  | M11 | c | 5 | b | ns | z
Extractable Tin | 2092-000122 Rev AA | 0 | HPLC grade Methylene Chloride | mL |  |  | M10 | c | 5 | b | ns | z
DHS | 2092-001026 Rev AB | 1 piece | Purified Nitrogen | 180mins x 65sccm |  |  | M7 | c | 5 | b | ns | z
GCMS - Hydrocarbon Residue | 2092-771690 Rev AD |  | HPLC grade n-Hexane | mL |  |  | M14 | c | 5 | b | ns | z
Results: | [5 empty] | 21 |  | r | a
The specification is based on Western Digital's document no.96-004575 Rev AU for Antidisk (3.5 Inch Enterprise) | [7 empty] | r | a
Required Test | Analytes | Specification Limits (ng/sqcm) | Results (ng/sqcm) | Method Detection Limit (ng/sqcm) | Pass / Fail |  | M1 | h | 6 | b | s | z
IC (Anions) | Fluoride as F | <10 | #DIV/0! | #DIV/0! | #DIV/0! |  | M1 | c | 6 | b | ns | z
| Chloride as Cl | <10 | #DIV/0! | #DIV/0! | #DIV/0! |  | M1 | c | 6 | b | ns | z
| Bromide as Br | <10 | #DIV/0! | #DIV/0! | #DIV/0! |  | M1 | c | 6 | b | ns | z
| Nitrate as NO3 | <10 | #DIV/0! | #DIV/0! | #DIV/0! |  | M1 | c | 6 | b | ns | z
| Phosphate as PO4 | <10 | #DIV/0! | #DIV/0! | #DIV/0! |  | M1 | c | 6 | b | ns | z
| Sulphate as SO4 | <10 | #DIV/0! | #DIV/0! | #DIV/0! |  | M1 | c | 6 | b | ns | z
| Total Anions | <30 | #DIV/0! | - | #DIV/0! |  | M1 | c | 6 | b | ns | z
Required Test | Analytes | Specification Limits (ng/sqcm) | Results (ng/sqcm) | Method Detection Limit (ng/sqcm) | Pass / Fail |  | M2 | h | 6 | b | s | z
IC (Cations) | Lithium as Li | NA | #DIV/0! | #DIV/0! | NA |  | M2 | c | 6 | b | ns | z
| Sodium as Na | NA | #DIV/0! | #DIV/0! | NA |  | M2 | c | 6 | b | ns | z
| Ammonium as NH4 | NA | #DIV/0! | #DIV/0! | NA |  | M2 | c | 6 | b | ns | z
| Potassium as K | NA | #DIV/0! | #DIV/0! | NA |  | M2 | c | 6 | b | ns | z
| Magnesium as Mg | NA | #DIV/0! | #DIV/0! | NA |  | M2 | c | 6 | b | ns | z
| Calcium as Ca | NA | #DIV/0! | #DIV/0! | NA |  | M2 | c | 6 | b | ns | z
| Total Cations | <30 | #DIV/0! | - | #DIV/0! |  | M2 | c | 6 | b | ns | z
Required Test | Analyte | Specification Limits (ng/sqcm) | Results (ng/sqcm) | Pass / Fail |  |  | M3M4 | h | 5 | b | s | z
NVR | Non-volatile residue (NVR) | <100 | #DIV/0! | #DIV/0! |  |  | M3M4 | c | 5 | b | ns | z
Required Test | Analyte | Specification Limits (ng/sqcm) | Results (ng/sqcm) | Pass / Fail |  |  | M3M5 | h | 5 | b | s | z
FTIR | Silicone @ wave nos. 2962, 1261, 1092, 1022 & 800cm-1 | <5 | Not Detected | #DIV/0! |  |  | M3M5 | c | 5 | b | ns | z
Required Test | Element | Specification Limits (NA) | Results (NA) | Method Detection Limit (NA) | Pass / Fail |  | M9 | h | 6 | b | s | z
Extractable Cu | Copper | NA | #DIV/0! | #DIV/0! | NA |  | M9 | c | 6 | b | ns | z
Required Test | Element | Specification Limits (NA) | Results (NA) | Method Detection Limit (NA) | Pass / Fail |  | M9 | h | 6 | b | s | z
Extractable Pt | Platinum | NA | #DIV/0! | #DIV/0! | NA |  | M9 | c | 6 | b | ns | z
Required Test | Element | Specification Limits (NA) | Results (NA) | Method Detection Limit (NA) | Pass / Fail |  | M10 | h | 6 | b | s | z
Extractable Sn | Tin | NA | #DIV/0! | #DIV/0! | NA |  | M10 | c | 6 | b | ns | z
Required Test | Analyte | Specification Limits (ng/part) | Results (ng/part) | Pass / Fail |  |  | M8 | h | 5 | b | s | z
CVR | Tributyl Tin Chloride | 0 | 0 | Pass |  |  | M8 | c | 5 | b | ns | z
Required Test | Particle Bin Size (um) | Specification Limits (particles/sqcm) | Average of 5 data points (particles/sqcm) | Pass / Fail |  |  | M6M26 | h | 5 | b | s | z
LPC | ≥ 0.3 | <3000 | 0 | Pass |  |  | M6 | c | 5 | b | ns | z
LPC | ≥ 0.5 | NA | 0 | NA |  |  | M6M26 | c | 5 | b | ns | z
Required Test | Analytes | Specification Limits (ng/part) | Results (ng/part) | Pass / Fail |  |  | M7 | h | 5 | b | s | z
DHS | Total Acrylate and Methacrylate | NA | Not Detected | NA |  |  | M7 | c | 5 | b | ns | z
| Total Siloxane | <30 | 4.9841 | PASS |  |  | M7 | c | 5 | b | ns | z
| - Cyclopentasiloxane, decamethyl- | - | 4.9841 |  |  |  | M7 | c | 5 | b | ns | z
| Total Unknown | - | 94.7511 |  |  |  | M7 | c | 5 | b | ns | z
| - Di-n-octylether-D34 | - | 0.42296 |  |  |  | M7 | c | 5 | b | ns | z
| - Octacosane | - | 4.6358 |  |  |  | M7 | c | 5 | b | ns | z
| - Docosane | - | 4.2971 |  |  |  | M7 | c | 5 | b | ns | z
| - Eicosane | - | 15.6382 |  |  |  | M7 | c | 5 | b | ns | z
| - Eicosane | - | 6.6673 |  |  |  | M7 | c | 5 | b | ns | z
| - Hexatriacontane | - | 3.2022 |  |  |  | M7 | c | 5 | b | ns | z
| - Heneicosane | - | 9.9491 |  |  |  | M7 | c | 5 | b | ns | z
| - Pentadecane, 8-heptyl- | - | 8.4677 |  |  |  | M7 | c | 5 | b | ns | z
| - Hexatriacontane | - | 3.1813 |  |  |  | M7 | c | 5 | b | ns | z
| - Tricyclo[4.2.1.1(2,5)]decane | - | 7.0036 |  |  |  | M7 | c | 5 | b | ns | z
| - 5-Allyl-2-norbornene | - | 2.5164 |  |  |  | M7 | c | 5 | b | ns | z
| - ENDO,EXO,ENDO-PERHYDRO-4,9 | - | 2.5577 |  |  |  | M7 | c | 5 | b | ns | z
| - Tetracyclo[7.1.0(1,6).0(4,9).0(8,1 | - | 17.9765 |  |  |  | M7 | c | 5 | b | ns | z
| - Tricyclo[5.2.1.0(2,6)]decan-8-ol | - | 5.347 |  |  |  | M7 | c | 5 | b | ns | z
| - Bicyclo[2.2.1]heptane, 2-(2-methyl | - | 2.8884 |  |  |  | M7 | c | 5 | b | ns | z
| Total Sulfur Compound | - | 2.7308 |  |  |  | M7 | c | 5 | b | ns | z
| - Benzothiazole | - | 2.7308 |  |  |  | M7 | c | 5 | b | ns | z
| Total Phthalate | - | Not Detected |  |  |  | M7 | c | 5 | b | ns | z
| Total Phenol | - | 18.3541 |  |  |  | M7 | c | 5 | b | ns | z
| - Phenol | - | 18.3541 |  |  |  | M7 | c | 5 | b | ns | z
| Total Others | - | 21.5663 |  |  |  | M7 | c | 5 | b | ns | z
| - 1,4-Dioxane | - | 2.2185 |  |  |  | M7 | c | 5 | b | ns | z
| - 1-Hexanol, 2-ethyl- | - | 2.7492 |  |  |  | M7 | c | 5 | b | ns | z
| - Benzyl Alcohol | - | 2.1125 |  |  |  | M7 | c | 5 | b | ns | z
| - Benzene, 1,4-dichloro- | - | 6.0959 |  |  |  | M7 | c | 5 | b | ns | z
| - Ethanol, 2-phenoxy- | - | 7.3615 |  |  |  | M7 | c | 5 | b | ns | z
| - CYCLOHEXADECAN, PERDEUTERO- | - | 1.0286 |  |  |  | M7 | c | 5 | b | ns | z
| Total Initiator | - | 17.7435 |  |  |  | M7 | c | 5 | b | ns | z
| - Benzaldehyde | - | 13.8667 |  |  |  | M7 | c | 5 | b | ns | z
| - Acetophenone | - | 3.8769 |  |  |  | M7 | c | 5 | b | ns | z
| Total Antioxidant | - | 30.8844 |  |  |  | M7 | c | 5 | b | ns | z
| - Phenol, 2,4-bis(1,1-dimethylethyl)- | - | 30.8844 |  |  |  | M7 | c | 5 | b | ns | z
| Total Hydrocarbon | - | 707.6941 |  |  |  | M7 | c | 5 | b | ns | z
| Total Aromatic Hydrocarbon | - | 115.9427 |  |  |  | M7 | c | 5 | b | ns | z
| - Toluene | - | 58.7737 |  |  |  | M7 | c | 5 | b | ns | z
| - Ethylbenzene | - | 11.5047 |  |  |  | M7 | c | 5 | b | ns | z
| - Benzene, 1,3-dimethyl- | - | 11.8219 |  |  |  | M7 | c | 5 | b | ns | z
| - Styrene | - | 1.6536 |  |  |  | M7 | c | 5 | b | ns | z
| - p-Xylene | - | 2.6096 |  |  |  | M7 | c | 5 | b | ns | z
| - Benzene, 1,3,5-trimethyl- | - | 7.5379 |  |  |  | M7 | c | 5 | b | ns | z
| - Naphthalene | - | 4.5862 |  |  |  | M7 | c | 5 | b | ns | z
| - 1,1'-Biphenyl, 2,2',5,5'-tetramethyl- | - | 6.5641 |  |  |  | M7 | c | 5 | b | ns | z
| - Naphthalene, 1,2,3-trimethyl-4-propenyl-, (E)- | - | 5.5932 |  |  |  | M7 | c | 5 | b | ns | z
| - 1,1-Bis(p-tolyl)ethane | - | 5.2979 |  |  |  | M7 | c | 5 | b | ns | z
| Total Aliphatic Hydrocarbon | - | 43.8291 |  |  |  | M7 | c | 5 | b | ns | z
| - Heptane | - | 1.0416 |  |  |  | M7 | c | 5 | b | ns | z
| - Nonane | - | 2.2686 |  |  |  | M7 | c | 5 | b | ns | z
| - Decane | - | 3.8509 |  |  |  | M7 | c | 5 | b | ns | z
| - Undecane | - | 3.5296 |  |  |  | M7 | c | 5 | b | ns | z
| - Dodecane | - | 3.8675 |  |  |  | M7 | c | 5 | b | ns | z
| - Tridecane | - | 1.4191 |  |  |  | M7 | c | 5 | b | ns | z
| - Tetradecane | - | 2.9763 |  |  |  | M7 | c | 5 | b | ns | z
| - Hexadecane | - | 2.5205 |  |  |  | M7 | c | 5 | b | ns | z
| - Octadecane | - | 1.8105 |  |  |  | M7 | c | 5 | b | ns | z
| - Hexadecane, 2-methyl- | - | 14.1644 |  |  |  | M7 | c | 5 | b | ns | z
| - Eicosane | - | 6.3802 |  |  |  | M7 | c | 5 | b | ns | z
| Hydrocarbon Hump | - | 547.9223 |  |  |  | M7 | c | 5 | b | ns | z
| - Hydrocarbon Hump(subtract) |  | 547.9223 |  |  |  | M7 | c | 5 | b | ns | z
| Total Outgassing | <2000 | 898.7085 | PASS |  |  | M7 | c | 5 | b | ns | z
Required Test | Analytes | Specification Limits (ng/part) | Result (ng/part) | Pass / Fail |  |  | M14 | h | 5 | b | s | z
GCMS | Total Hydrocarbon between C18-C40 | <200 |  | PASS |  |  | M14 | c | 5 | b | ns | z
| - Total Hydrocarbon Peak between C18-C40 | - |  |  |  |  | M14 | c | 5 | b | ns | z
| - Total Hydrocarbon Hump between C18-C40 | - |  |  |  |  | M14 | c | 5 | b | ns | z
| - Total Hydrocarbon Hump between C18-C40 + DP154/ DP155 | - |  |  |  |  | M14 | c | 5 | b | ns | z
| Surfactant | - |  |  |  |  | M14 | c | 5 | b | ns | z
| - Triton X100 | - |  |  |  |  | M14 | c | 5 | b | ns | z
| - DP154/ DP155 | - |  |  |  |  | M14 | c | 5 | b | ns | z
| Fatty Acid Ester | <50 |  | PASS |  |  | M14 | c | 5 | b | ns | z
| - Tetradecanoic acid, hexadecyl ester | - |  |  |  |  | M14 | c | 5 | b | ns | z
| - Hexadecanoic acid, hexadecyl ester | - |  |  |  |  | M14 | c | 5 | b | ns | z
| - Octadecanoic acid, hexadecyl ester | - |  |  |  |  | M14 | c | 5 | b | ns | z
| Irgafos and Derivatives | <50 |  | PASS |  |  | M14 | c | 5 | b | ns | z
| - Irgafos | - |  |  |  |  | M14 | c | 5 | b | ns | z
| - Irgafos-oxidized | - |  |  |  |  | M14 | c | 5 | b | ns | z
| Total Organic Compound | - |  |  |  |  | M14 | c | 5 | b | ns | z
| Major Compounds (R.T) | Classification | Result (ng/part) |  |  |  | M14 | h | 5 | b | ns | z
[7 empty] | M14 | c | 5 | b | ns | z
[7 empty] | M14 | c | 5 | b | ns | z
[7 empty] | M14 | c | 5 | b | ns | z
[7 empty] | M14 | c | 5 | b | ns | z
[7 empty] | M14 | c | 5 | b | ns | z
DHS Chromatogram | [6 empty] | M7 | r | a
[7 empty] | M7 | r | a
[7 empty] | M14 | r | a
HC Chromatogram | [6 empty] | M14 | r | a
[7 empty] | M14 | r | a
[7 empty] | M3M5 | r | a
FTIR Spectrum | [6 empty] | M3M5 | r | a
[7 empty] | M3M5 | r | a
[7 empty] | M3M6 | r | a
Remarks: Surfactant compound was reported based on customer's request. | [6 empty] | M14 | r | a
| Tested By: | LS/YK | Approved By :
| Date: | 2013-03-01 | Date :
Worksheet for Western Digital's Liquid Particle Count | [6 empty] | M6M26 | r | a
M6M26
OUR REF NO. : | R/ATT/ELP/13/0532-D | [5 empty] | M6M26 | h | 2 | nb | ns | z
COMPANY : | Enpro Products (Thailand) Co., Ltd. | [5 empty] | M6M26 | c | 2 | nb | ns | z
DATE SAMPLE RECEIVED: | 2013-02-14 | [5 empty] | M6M26 | c | 2 | nb | ns | z
SAMPLE INFO: | Description: Antitle Disk | [5 empty] | M6M26 | c | 2 | nb | ns | z
MODEL: | Model: Tahoe | [5 empty] | M6M26 | c | 2 | nb | ns | z
PART NO.: | Remark: 24 pcs | [5 empty] | M6M26 | c | 2 | nb | ns | z
DATE ANALYSED: | 2013-02-28 | [5 empty] | M6M26 | c | 2 | nb | ns | z
OPERATOR: | 0 | [5 empty] | M6M26 | c | 2 | nb | ns | z
M6M26
M6M26
Test Method: 92-004230 Rev AJ 92-004134 Rev CD (For Tray Use) | [6 empty] | M6M26 | h | 2 | b | s | a
Surface Area, cm² | 1 | [5 empty] | M6M26 | c | 2 | b | ns | z
Extraction Vol, ml | 1 | [5 empty] | M6M26 | c | 2 | b | ns | z
Sample Vol, ml | 10 | [5 empty] | M6M26 | c | 2 | b | ns | z
No. of Parts | 1 | [5 empty] | M6M26 | c | 2 | b | ns | z
M6M26
Tank Conditions | Tank Volume | Sonication Freq. | Ultrasonic Power |  |  |  | M6M26 | h | 4 | b | s | z
| 20.2 L | 68 kHz | 4.8 W/L |  |  |  | M6M26 | c | 4 | b | ns | z
M6M26
Run | Accumulative Size (µm) | Blank (Counts/ml) | Sample (Counts/ml) | Blank-corrected (Counts/part) | Blank-corrected (Counts/cm²) |  | M6 | h | 6 | b | s | z
1 | 0.3 | 0 | 0 | 0 | 0 |  | M6 | c | 6 | b | ns | z
| 0.5 | 0 | 0 | 0 | 0 |  | M6 | c | 6 | b | ns | z
| 0.7 | 0 | 0 | 0 | 0 |  | M6M26 | c | 6 | b | ns | z
| 1 | 0 | 0 | 0 | 0 |  | M6M26 | c | 6 | b | ns | z
| 2 | 0 | 0 | 0 | 0 |  | M6M26 | c | 6 | b | ns | z
[7 empty] | M6 | c | 6 | b | s | a
2 | 0.3 | 0 | 0 | 0 | 0 |  | M6 | c | 6 | b | ns | z
| 0.5 | 0 | 0 | 0 | 0 |  | M6 | c | 6 | b | ns | z
| 0.7 | 0 | 0 | 0 | 0 |  | M6M26 | c | 6 | b | ns | z
| 1 | 0 | 0 | 0 | 0 |  | M6M26 | c | 6 | b | ns | z
| 2 | 0 | 0 | 0 | 0 |  | M6M26 | c | 6 | b | ns | z
[7 empty] | M6 | c | 6 | b | s | a
3 | 0.3 | 0 | 0 | 0 | 0 |  | M6 | c | 6 | b | ns | z
| 0.5 | 0 | 0 | 0 | 0 |  | M6 | c | 6 | b | ns | z
| 0.7 | 0 | 0 | 0 | 0 |  | M6M26 | c | 6 | b | ns | z
| 1 | 0 | 0 | 0 | 0 |  | M6M26 | c | 6 | b | ns | z
| 2 | 0 | 0 | 0 | 0 |  | M6M26 | c | 6 | b | ns | z
[7 empty] | M6M26 | b
Worksheet for Western Digital's Liquid Particle Count | [6 empty] | M6M26 | r | a
R/ATT/ELP/13/0532-D | [6 empty] | M6M26 | r | a
M6M26
M6M26
Run | Accumulative Size (µm) | Blank (Counts/ml) | Sample (Counts/ml) | Blank-corrected (Counts/part) | Blank-corrected (Counts/cm²) |  | M6 | h | 6 | b | s | z
4 | 0.3 | 0 | 0 | 0 | 0 |  | M6 | c | 6 | b | ns | z
| 0.5 | 0 | 0 | 0 | 0 |  | M6 | c | 6 | b | ns | z
| 0.7 | 0 | 0 | 0 | 0 |  | M6M26 | c | 6 | b | ns | z
| 1 | 0 | 0 | 0 | 0 |  | M6M26 | c | 6 | b | ns | z
| 2 | 0 | 0 | 0 | 0 |  | M6M26 | c | 6 | b | ns | z
[7 empty] | M6M26 | c | 6 | b | s | a
5 | 0.3 | 0 | 0 | 0 | 0 |  | M6 | c | 6 | b | ns | z
| 0.5 | 0 | 0 | 0 | 0 |  | M6 | c | 6 | b | ns | z
| 0.7 | 0 | 0 | 0 | 0 |  | M6M26 | c | 6 | b | ns | z
| 1 | 0 | 0 | 0 | 0 |  | M6M26 | c | 6 | b | ns | z
| 2 | 0 | 0 | 0 | 0 |  | M6M26 | c | 6 | b | ns | z
M6
Statistics | Accumulative Size (µm) | Blank-corrected (Counts/part) | Blank-corrected (Counts/cm²) |  |  |  | M6 | h | 4 | b | s | z
Average | 0.3 | 0 | 0 |  |  |  | M6 | c | 4 | b | ns | z
| 0.5 | 0 | 0 |  |  |  | M6 | c | 4 | b | ns | z
| 0.7 | 0 | 0 |  |  |  | M6M26 | c | 4 | b | ns | z
| 1 | 0 | 0 |  |  |  | M6M26 | c | 4 | b | ns | z
| 2 | 0 | 0 |  |  |  | M6M26 | c | 4 | b | ns | z
[7 empty] | M6 | c | 4 | b | s | a
Standard Deviation | 0.3 | 0 | 0 |  |  |  | M6 | c | 4 | b | ns | z
| 0.5 | 0 | 0 |  |  |  | M6 | c | 4 | b | ns | z
| 0.7 | 0 | 0 |  |  |  | M6M26 | c | 4 | b | ns | z
| 1 | 0 | 0 |  |  |  | M6M26 | c | 4 | b | ns | z
| 2 | 0 | 0 |  |  |  | M6M26 | c | 4 | b | ns | z
[7 empty] | M6 | c | 4 | b | s | a
%RSD Deviation | 0.3 | #DIV/0! | #DIV/0! |  |  |  | M6 | c | 4 | b | ns | z
| 0.5 | #DIV/0! | #DIV/0! |  |  |  | M6 | c | 4 | b | ns | z
| 0.7 | #DIV/0! | #DIV/0! |  |  |  | M6M26 | c | 4 | b | ns | z
| 1 | #DIV/0! | #DIV/0! |  |  |  | M6M26 | c | 4 | b | ns | z
| 2 | #DIV/0! | #DIV/0! |  |  |  | M6M26 | c | 4 | b | ns | z
[7 empty] | EOF | eof
eof
Version: | WD1.4
Date Created: | 27 July 2010 for Singapore
Formulas:
R/ATT/ELP/13/0532-D = [1]LPC!B6
ชนิดการทดสอบ, วิธีการทดสอบที่ใช้จำนวนตัวอย่างที่ใช้
สเปคของลูกค้าที่เลือก ขึ้นกับชนิดตัวอย่าง